## Slide 2: Organic Electrochemical Transistors (OECT) combined with enzyme glucose dehydrogenase as glucose sensors
Emmanuel Joseph Shaji
Bremen 11/11/2025

## Slide 3: Motivation Slide
Glucose is essential carbohydrate that serves as a primary energy source for most living organisms. It helps in:-
Diagnosis :- Diabetes, hyperglycemia and hypoglycemia.
Health Monitoring :- Physical energy and Cognitive function.
Masters Project Emmanuel Joseph Shaji
11/11/2025
1
Source: Google images

## Slide 4: Motivation Slide
Masters Project Emmanuel Joseph Shaji
11/11/2025
2
Non invasive
Expensive
Non invasive
No real time detection
Non invasive
Complex circuitry
High sensitivity
Invasive
Notes: Optical glucose sensing: Detects glucose through changes in light absorption, fluorescence, or scattering caused by glucose–enzyme or glucose–indicator interactions.
Urine/Saliva glucose testing: Determines glucose concentration by measuring the color change of a dye after enzymatic oxidation, compared to an optical or color reference.
Microwave/Impedance glucose sensing: Senses glucose by monitoring variations in dielectric properties or impedance of the sample as glucose concentration alters its conductivity and permittivity.
Electrochemical glucose sensing: Measures electrical current or potential generated by the enzymatic oxidation of glucose, proportional to its concentration.

## Slide 5: New candidate platform OECT
Working Principle
Masters Project Emmanuel Joseph Shaji
11/11/2025
3
Low power consumption
Real time monitoring
Low cost
Advantages
Flexibility and biocompatibility
Reduced invasiveness
Notes: . Ion injection from an electrolyte into an organic semiconductor channel to modulate its conductivity, thereby controlling the drain current in response to a gate voltage.

## Slide 6: GOx-OECT
GOx
What is GOx?
Challenges
GOx
Masters Project Emmanuel Joseph Shaji
11/11/2025
4
  Reproducibility Issues
  Stability Concerns
  Oxygen Dependence
  Hydrogen Peroxide Byproduct
Notes: .Widely used enzyme for glucose sensing-GOx (Glucose Oxidase) is an enzyme that specifically oxidizes glucose to produce H₂O₂, which generates a measurable electrical signal proportional to glucose concentration

## Slide 7: GDH- an alternative for GOx?
What is GDH?
No H2O2 By-Product
25 times higher activity than GOx
No oxygen needed
Masters Project Emmanuel Joseph Shaji
11/11/2025
5
Advantages
Notes: . GDH (Glucose Dehydrogenase) is an enzyme that catalyzes the oxidation of glucose using a cofactor (like PQQ), producing electrons that generate a measurable electrical signal proportional to glucose concentration.

## Slide 8: Research Question
Can Organic Electrochemical Transistors (OECTs) combined with the enzyme glucose dehydrogenase (GDH) be used to perform stable and reproducible glucose sensing?
Masters Project Emmanuel Joseph Shaji
11/11/2025
6

## Slide 9: Preparation of OECT for Testing
mPEDOT:PSS
mPEDOT:PSS
Pure PEDOT:PSS
Spincoat
(1000rpm for 60s)
Spincoat
(1000rpm for 60s)
Bake
(110°C for 10 min)
Bake
(110°C for 10 min)
PEDOT:PSS
EG
DBSA
GOPS
mPEDOT:PSS
Masters Project Emmanuel Joseph Shaji
11/11/2025
7
Notes: External Platinum gate was used because none of the chips had inbuilt Platinum in it. Why I used dual layer channel will be explained later

## Slide 10: Enzyme Preparation
PBS
Buffer
GOx
Enzyme
PBS
Buffer
PBS
Buffer
GOx in PBS
Buffer
PBS
Buffer
GDH
Enzyme
MES
Buffer
PBS
Buffer
GDH-PQQ in MES
Buffer
CaCl2
NaOH
GOx
GDH
pH 7
Masters Project Emmanuel Joseph Shaji
11/11/2025
8
PQQ
Notes: . They help hold PQQ firmly in the active site, allowing proper electron transfer from glucose to PQQ.
Without Ca²⁺, PQQ can detach or remain inactive, leading to poor or no enzymatic activity.
PQQ (Pyrroloquinoline Quinone) is the redox cofactor (electron carrier) of PQQ–GDH.
It participates directly in the oxidation of glucose to gluconolactone by accepting electrons and protons from glucose.
Enzyme Cofactor Preferred Buffer Reason GDH (PQQ–GDH) PQQ + Ca²⁺
MES Prevents Ca²⁺ precipitation, maintains cofactor stability GOx FAD
PBS Maintains physiological pH, compatible with FAD and O₂ reaction

## Slide 11: Experiments done and Glucose Samples tested
Control
Enzyme immobilized
Steady state
Floating Enzyme
Steady state
Floating Enzyme
Real time
1nM | 2nM | 3nM | 4nM | 5nM
1nM | 10nM | 100nM | 1µM | 10µM | 100µM | 1mM | 10mM
1µM | 10µM | 100µM | 1mM | 10mM
No Glucose
Masters Project Emmanuel Joseph Shaji
11/11/2025
9
Notes: 1nM-5nM (Wildau‘s Strategy)
1nM-10mM ( to learn more intermediate concentrations)
1µM-10mM ( Small concentrations didnt show sensing)

## Slide 12: Control Experiment
Vg- 0 V to 0.6 V at a scan rate of 0.01 V/s and Vd constant at –0.2 V with external overhanging Pt gate.
Masters Project Emmanuel Joseph Shaji
11/11/2025
10

## Slide 13: Steady state measurement of OECT with GDH enzyme immobilized on channel
Vg- 0 V to 0.6 V at a scan rate of 0.01 V/s and Vd constant at –0.2 V with external overhanging Pt gate.
Transfer characteristics
Max transconductance vs glucose concentration
Masters Project Emmanuel Joseph Shaji
11/11/2025
11
Notes:     On current (Ion)
    Off current (Ioff)
    On/Off ratio (Ion / Ioff)
    Maximum transconductance (gm)
    Average current modulation (|Ioff – Ion| / Ion)

## Slide 14: Steady state measurement of OECT with GDH enzyme immobilized on channel
On/Off Ratio vs glucose concentration
Avg current modulation vs glucose concentration
Masters Project Emmanuel Joseph Shaji
11/11/2025
12

## Slide 15: Steady state measurement of OECT with GOx enzyme as floating electrolyte
Vg- 0 V to 0.6 V at a scan rate of 0.01 V/s and Vd constant at –0.2 V with external overhanging Pt gate.
Transfer characteristics
Max transconductance vs glucose concentration
Masters Project Emmanuel Joseph Shaji
11/11/2025
13
Notes: OECts are sensitive to tiny ionic changes, ions starts to penetrate the channel and reduce its oxidized state(PEDOT+ ---- PEDOT0) as gate voltage is applied. The increase in glucose concentration increase ions

## Slide 16: Steady state measurement of OECT with GOx enzyme as floating electrolyte
On/Off Ratio vs glucose concentration
Avg current modulation vs glucose concentration
Masters Project Emmanuel Joseph Shaji
11/11/2025
14

## Slide 17: Steady state measurement of OECT with GDH enzyme as floating electrolyte
Vg- 0 V to 0.6 V at a scan rate of 0.01 V/s and Vd constant at –0.2 V with external overhanging Pt gate.
Transfer characteristics
Max transconductance vs glucose concentration
Masters Project Emmanuel Joseph Shaji
11/11/2025
15

## Slide 18: Steady state measurement of OECT with GDH enzyme as floating electrolyte
Avg current modulation vs glucose concentration
On/Off Ratio vs glucose concentration
Masters Project Emmanuel Joseph Shaji
11/11/2025
16

## Slide 19: Real time measurement of OECT with GOx enzyme as floating electrolyte
Id vs time
ΔId vs Glucose concentration
Measurement setup: Drain voltage fixed at Vd = –0.2 V and gate voltage held at Vg = 0.6 V with external overhanging Pt gate.
Masters Project Emmanuel Joseph Shaji
11/11/2025
17
Notes:   Id vs Time: Shows real-time current change after each glucose addition.
  ΔId vs Glucose Concentration: Shows overall current variation across glucose levels

## Slide 20: Real time measurement of OECT with GDH enzyme as floating electrolyte
ΔId vs Glucose concentration
Id vs time
Measurement setup: Drain voltage fixed at Vd = –0.2 V and gate voltage held at Vg = 0.6 V with external overhanging Pt gate.
Masters Project Emmanuel Joseph Shaji
11/11/2025
18

## Slide 21: Conclusion
19
In control, dual-layer PEDOT:PSS channels (mPEDOT:PSS + PH1000) showed higher stability and reproducibility than single-layer devices.
Floating-enzyme configuration provided consistent signals, while immobilized enzymes tended to detach.
GOx-based OECTs: qualitative detection (1 µM–100 µM, steady-state) and linear response (10 µM–1 mM, real-time) — ideal for saliva and sweat sensing.
GDH-based OECTs: qualitative detection (100 nM–10 µM, steady-state) and linearity (10 µM–10 mM, real-time) — covering both biofluid and blood glucose levels.
Enzyme type and configuration strongly affect OECT glucose sensing performance and stability.
Masters Project Emmanuel Joseph Shaji
11/11/2025

## Slide 22
19
Lower gate voltage from 0.6 V to 0.3 V for reaching saturation faster in real-time signals.
Immobilize enzyme on the gate electrode to improve stability and sensitivity.
Masters Project Emmanuel Joseph Shaji
11/11/2025
Future Perspectives

## Slide 23: Questions?
20
Masters Project Emmanuel Joseph Shaji
11/11/2025

## Slide 24
Masters Project Emmanuel Joseph Shaji
11/11/2025

## Slide 25
Masters Project Emmanuel Joseph Shaji
11/11/2025

## Slide 26: References
[1] S. Ferri, K. Kojima, and K. Sode, “Review of glucose oxidases and glucose dehydrogenases: a bird's eye view of glucose sensing enzymes,” J. Diabetes Sci. Technol., vol. 5, no. 5, pp. 1068–1076, 2011, doi: 10.1177/193229681100500507.
[2] E. M. Thaning, M. L. M. Asplund, T. A. Nyberg, O. W. Inganäs, and H. von Holst, “Stability of poly(3,4-ethylene dioxythiophene) materials intended for implants,” J. Biomed. Mater. Res. Part B Appl. Biomater., vol. 94B, no. 1, pp. 254–262, 2010, doi: 10.1002/jbm.b.31597.
[3] C. O. Malikkides and R. H. Weiland, “On the mechanism of immobilized glucose oxidase deactivation by hydrogen peroxide,” Biotechnol. Bioeng., vol. 74, no. 1, pp. 109–116, 2001, doi: 10.1002/bit.260241109.

## Slide 27
Quantity | Low glucose | Medium glucose | High glucose (saturation)
I_ON | High | Decreasing | Constant (saturated)
I_OFF | Low | Decreasing | Constant (saturated)
Transconductance | High | Decreasing | Plateau / minimum
On–off ratio | Moderate | Increasing | Saturated / plateau
Current modulation | Moderate | Increasing | Saturated / plateau

## Slide 28: Appendix
22
Masters Project Emmanuel Joseph Shaji
11/11/2025
1.Steady state measurement of OECT with GDH enzyme immobilized on channel
OECT 2
OECT 3

## Slide 29: Appendix
Masters Project Emmanuel Joseph Shaji
11/11/2025
2.Steady state measurement of OECT with GOx enzyme as floating electrolyte
OECT 2
OECT 3
23

## Slide 30: Appendix
24
Masters Project Emmanuel Joseph Shaji
11/11/2025
3.Steady state measurement of OECT with GDH enzyme as floating electrolyte
OECT 2
OECT 3

## Slide 31: Appendix
25
Masters Project Emmanuel Joseph Shaji
11/11/2025
3.Steady state measurement of OECT with GDH enzyme as floating electrolyte
OECT 4

## Slide 32: Appendix
26
Masters Project Emmanuel Joseph Shaji
11/11/2025
OECT 2
OECT 3
4.Real time measurement of OECT with GOx enzyme as floating electrolyte

## Slide 33: Appendix
Masters Project Emmanuel Joseph Shaji
11/11/2025
5.Real time measurement of OECT with GDH enzyme as floating electrolyte
OECT 2
OECT 3
27

## Slide 34: 6.Fabrication of OECT
Mask Design and Layout
  Chip layout (21 × 21 mm) with 3 channels (500 × 500 µm) designed in AutoCAD and aligned on wafer template
2. Electrode Deposition and Patterning
  Gold source, drain, and gate interconnects patterned on Si/SiO₂ wafer using photolithography and PI insulation layers
3. PEDOT:PSS Deposition
  Modified PEDOT:PSS (with EG, DBSA, GOPS) spin-coated, baked at 110 °C, similarly pure PEDOT:PSS on top and sacrificial PI layer lifted off to form final channel.
28
Masters Project Emmanuel Joseph Shaji
05/11/2025
Appendix
Notes: 1.Mask Design and LayoutDesigned electrode and channel geometries in AutoCAD (21 mm × 21 mm chip)Included three 500 μm × 500 μm channels and in-plane Pt gates (GT1–GT3) Overhanging platinum-coated glass electrode used as gate during testing Layout exported in .dxf format and aligned using wafer template(Each chip design is aligned by snapping its bottom-left corner to the top-right corner of a positioning square on the wafer template.)
2. Electrode Deposition and Patterning-Started with thermally oxidized Si wafer (SiO₂ ≈ 500 nm) for insulation. Surface treated with APTES → improved adhesion for polyimide (PI)First PI layer spin-coated and cured as adhesion/buffer layer. Gold film sputtered and patterned (source, drain, gate lines) via first photolithography, Second- Third PI layer added as sacrificial layer for later lift-off Second photoresist (PR2) applied to define contact vias and channels. Reactive-ion etching (RIE) used to open via holes and define channels. PR2 removed after etching.
3. PEDOT:PSS Deposition Formulated modified PEDOT:PSS (mPEDOT:PSS):PEDOT:PSS (Clevios PH1000) + 5% EG + 0.1% DBSA + 1% GOPS Mixed and stirred for 15 minutes Spin-coated (1000 rpm, 1 min) onto exposed channel region Baked at 110 °C for 10 min for solvent removal Sacrificial PI layer lifted off, revealing final channel geometry

## Slide 35: 7.Enzyme Preparation
29
GDH
GOx
Masters Project Emmanuel Joseph Shaji
05/11/2025
Appendix
Notes: Glucose Oxidase specific activity=360U/mg
Unit/Volume needed=500 U/ml
Required mass=specific activity/Unit/Volume=1.39mg/ml
Volume prepared=5mL
Unit/Volume Prepared=5ml*1.39mg/ml=6.95mg

## Slide 36: 8.Glucose Preparation
30
PBS
(GOx)
MES
(GDH)
MES
(GDH)
Masters Project Emmanuel Joseph Shaji
05/11/2025
Appendix

## Slide 37: Motivation Slide
Masters Project Emmanuel Joseph Shaji
11/11/2025
2
Non invasive
Real time detection
Light scattering & Tissue absorption
Expensive
Non invasive
Simple design
Indirect measurement
No real time detection
Non invasive
Real time detection
Complex circuitry
Still in study phase
Well established
High sensitivity
Invasive
Temparature and pH dependent
Notes: Optical glucose sensing: Detects glucose through changes in light absorption, fluorescence, or scattering caused by glucose–enzyme or glucose–indicator interactions.
Urine/Saliva glucose testing: Determines glucose concentration by measuring the color change of a dye after enzymatic oxidation, compared to an optical or color reference.
Microwave/Impedance glucose sensing: Senses glucose by monitoring variations in dielectric properties or impedance of the sample as glucose concentration alters its conductivity and permittivity.
Electrochemical glucose sensing: Measures electrical current or potential generated by the enzymatic oxidation of glucose, proportional to its concentration.

## Slide 38
Substrate
Source
Drain
mPEDOT:PSS
Vds
Vgs
Electrolyte
Gate
Electrolyte + GOx
Pure PEDOT:PSS

## Slide 39
Substrate
Source
Drain
mPEDOT:PSS
Electrolyte
Gate
Pure PEDOT:PSS
GDH
Pure PEDOT:PSS
Au
GDH
mPEDOT:PSS
Au
GDH
Au
mPEDOT:PSS
Pure PEDOT:PSS
a)
b)
c)

## Slide 40: OECT fabrication with 3 PI layers
Universität Bremen
IMSAS
Institute for Microsensors,
-actuators and -systems
Faculty 01
Physics / Electrical Engineering
Henrique F. P. Barbosa
Bremen 08.01.2024

## Slide 41
SiO2
SiO2
PI
Au
SiO2
PI
Au
PR1
SiO2
PI
Au
PR1
SiO2
PI
Au
SiO2
PI
Au
SiO2
PI
Au
SiO2
PI
MA
Mask
PR1
Mask
PR1
Wafer
Au sputtering over PI
PR1 deposition
1st Mask exposure
Post-1st exposure
Mask removal
PR1 development
Au etching
Si
Si
Si
Si
Si
Si
Si
Si
APTES: adhesion promoter

## Slide 42
MA
PI
SiO2
PI
PR2
Short RIE exposure
PI double layer deposition
PR2 deposition
2nd Mask exposure
Post-2nd exposure
Mask removal
PR2 development
PR2
Mask
PR2
PR2
PR2
Mask
SiO2
PI
PR1 removal
SiO2
PI
PI
PI
SiO2
PI
PI
RIE
PI
SiO2
PI
PI
PI
SiO2
PI
PI
PI
SiO2
PI
PI
PI
SiO2
PI
PI
Si
Si
Si
Si
Si
Si
Si
Si

## Slide 43
PI
SiO2
PI
PI
Ag
RIE
Post-RIE
RIE
PR2 removal
PI top layer peeling
(structuring)
PR2
PR2
OSC deposition
Ag
PI
SiO2
PI
PI
PI
SiO2
PI
PI
PI
SiO2
PI
PI
OSC
PI
SiO2
PI
Si
Si
Si
Si
Si
PI
SiO2
PI
PI
Si
LAYER SIZES
200 nm
500 nm
4.7 µm
4.7 µm
300 nm